[image: image1.png]meridian BIOSCIENCE
S





Page 3 of 8

Procedure:  Alethia C. difficile
Procedure and revision number: SN11016 Rev. 01/20

Meridian Catalog #480050
	Institution: 

	Address:

	Department:


	Prepared By:
	Date Adopted
	Supersedes procedure #:

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	


	Distributed To:
	# of Copies
	Distributed To:
	# of Copies

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	


NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might affect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.
PRINCIPLE:
The Alethia C. difficile DNA amplification assay, performed on the Alethia reader, is a qualitative in vitro diagnostic test for the direct detection of toxigenic C. difficile in human stool specimens from pediatric and adult patients suspected of having Clostridium difficile-associated disease (CDAD).

The Alethia C. difficile assay utilizes loop-mediated isothermal DNA amplification (LAMP) technology to detect the pathogenicity locus (PaLoc) of toxigenic Clostridium difficile.  The Clostridium difficile PaLoc is a gene segment present in all known toxigenic C. difficile strains.  The C. difficile PaLoc codes for both the Toxin A gene (tcdA) and the Toxin B gene (tcdB), has conserved border regions, and is found at the same site on the C. difficile genome for all toxigenic strains.  The Alethia C. difficile assay detects the PaLoc by targeting a partial DNA fragment on the Toxin A gene.  The tcdA target region was selected as an intact region remaining in all known A+B+ and A-B+ toxinotypes.

Alethia C. difficile is intended for use in hospital, reference or state laboratory settings.  The device is not intended for nonlaboratory point-of-care use.

Toxigenic Clostridium difficile is a major cause of antibiotic associated diarrhea and colitis and is the causative agent for virtually all cases of pseudomembranous colitis.  Although about 2% of normal healthy adults are colonized with C. difficile, many patients acquire this organism through nosocomial infection.  Exposure to most antibiotics is thought to allow proliferation of toxigenic C. difficile by disrupting the normal intestinal flora. Two large toxin proteins (TcdA [or toxin A] and TcdB [toxin B]) are thought to be the primary virulence factors of C. difficile.  These toxins are encoded by two separate genes, named tcdA and tcdB, respectively.  Together, with three additional genes, they form a 19.6kb pathogenicity locus called PaLoc.  
SPECIMEN:
Preferred Sample Types:

1.   Human stool samples, unpreserved
2.   Human stool samples, preserved in Cary-Blair-based media
3.   Unformed samples indicative of CDAD
Undesirable samples: Formed stools with no indication of CDAD, colonic aspirates
Collection:

Human stool samples, unpreserved: Samples should be transported and stored at 2-8 C prior to testing.  Samples should be tested as soon as possible, but may be held up to 24 hours at 21-27 C or 5 days at 2-8 C.  
Human stool samples, preserved in Cary-Blair-based media: Samples should be transported and stored at 2-8 C prior to testing.  Samples should be tested as soon as possible, but may be held up to 5 days at 2-8 C.  
This facility’s procedure for specimen collection is: _____________________________________ ______________________________________________________________________________
This facility’s procedure for transporting specimens is: __________________________________ ______________________________________________________________________________
This facility’s procedure for rejected specimens is: _____________________________________ ______________________________________________________________________________
MATERIALS AND EQUIPMENT:
MATERIALS:
1. Alethia C.difficile Sample Preparation Apparatus:  Sampling unit consisting of sample preparation chamber, dropper tip, cap and Sample Dilution Buffer (Phosphate Buffered Saline and formalin treated Staphylococcus aureus, with Sodium Azide (0.09%)as a preservative) 
2. Alethia Reaction Buffer: Tris-buffered solution containing sodium azide (0.09%)as a preservative
3. Alethia C.difficile Test Device: Two separate chambers containing dry reagent lyospheres comprised of DNA polymerase, Deoxyribonucleoside Triphosphate (dNTPs), and either C. difficile specific primers (TEST Chamber) or S. aureus primers (CONTROL Chamber)
4. Sample Collection Brushes
5. Alethia Heat Treatment Tubes
MATERIALS PROVIDED SEPARATELY

Alethia External Control Kit, Catalog Number: 479920

MATERIALS NOT PROVIDED

1. Disposable latex gloves, powder free

2. DNAse/RNAse free, aerosol resistant pipette tips

EQUIPMENT:

1. Dry-bath with 12mm heat block capable of 95 C
2.  Digital Thermometer with Max/Min Temperature Memory (e.g. Traceable® Lollipop™ Waterproof/Shockproof Thermometer)
3. Vortex mixer
4.  Interval timer
5.  Micropipette capable of dispensing 50 μL
6.  Micropipette capable of dispensing 200 μL
7.  Alethia reader 
PERFORMANCE CONSIDERATIONS:
1.
All reagents are for in vitro diagnostic use only.
2. Follow Biosafety Level 2 and Good Laboratory Practices during testing.  Treat all specimens and used test devices as capable of transmitting infectious disease.  Do not eat, drink or smoke in areas where specimens or kit reagents are handled.
3. Wear disposable gloves while handling specimens and thoroughly wash hands afterwards.
4. Quality Control Programs for Molecular Testing Laboratories should be employed.
5. Alethia Sample Dilution Buffer contains formalin inactivated organisms and should be treated as potentially infectious.
6. The Alethia C. difficile Test Device contains lyophilized reagents.  The protective pouch should not be open until ready to perform the assay.
7. The Alethia C. difficile Test Device includes a latch feature that is designed to prevent contamination of the test area with amplification product. Do NOT use Test Devices with broken latches.
8. Dispose of used Alethia Test Devices immediately after processing, leaving the device latch securely in place. Opening the device after amplification may result in contamination of the test area with amplification product. 
SHELF LIFE AND STORAGE:
1. The expiration date is indicated on the kit label.  
2. Store the kit at 2 - 27 C.
At this facility, kits are stored:_______________________________________________
CALIBRATION: 
Not applicable to this assay.
QUALITY CONTROL:
This test should be performed by qualified personnel per local regulatory requirements.
1. Each device contains an internal control well that controls for amplification inhibition, assay reagents and sample processing effectiveness.
2. The heat treatment step is monitored with an external thermometer and interval timer.  Use the Max/Min temperature memory of the thermometer to ensure that a temperature of 95 +/- 5 C is maintained.  Use the interval timer to ensure that heat-treatment duration is 10 +/- 2 minutes.  
3. Good laboratory practice recommends the use of control materials.  Users should follow the appropriate federal, state and local guidelines concerning the running of external quality controls. 

4. Alethia C. difficile External Control Reagents are supplied separately (Catalog 479920).  It is recommended that the reactivity of each new lot and each new shipment of Alethia C. difficile be verified on receipt and before use.  External control tests should be performed in accordance with appropriate federal, state, and local guidelines.  The Alethia C. difficile test kit should not be used in patient testing if the external controls do not produce the correct results.

5. A separate device must be used for each external control.  

6. DO NOT heat-treat the Positive or Negative Control samples.

QC Testing Frequency and Documentation:
For this facility, External QC is run: __________________________________________________                                                                                                   

Results of External QC and action(s) taken when control results are unacceptable are documented: ___________________________________________________________________
REAGENT PREPARATION:
Ensure kit reagents are at room temperature (21 - 27 C) before use.  Incorrect results may be obtained if reagents are not brought to room temperature prior to use.
PROCEDURE:
NOTE:  Ensure that the Alethia instrument is powered on and required performance verifications have been completed prior to initiation of SPECIMEN PREPARATION. Refer to the Alethia reader Operator’s Manual for further information regarding instrument set-up and operation. 

SPECIMEN PREPARATION:
1.
Mix stool sample thoroughly. 
2.
Collect mixed sample specimen using Sample Collection Brush.

a. Liquid Stool:  Immerse Sample Collection Brush completely into specimen.

b. Semi-solid Stool: Rotate Sample Collection Brush over specimen surface, lightly coating the surfaces of the brush bristles. Approximately one-half of the brush should be coated; over-collection of stool may lead to clogging of the sample collection device. 
3.
Add the Sample Collection Brush to the Alethia C. difficile Sample Preparation Apparatus containing Sample Diluent. Break-off the Sample Collection Brush at the ‘Break’ mark on the handle.  Replace and secure the Sample Preparation Apparatus cap.  For stool in Cary-Blair media, transfer 200 µL of specimen to the Alethia C. difficile Sample Preparation Apparatus and secure the cap.  Vortex the Sample Preparation Apparatus for a minimum of 10 seconds.  Specimen in Alethia C. difficile Sample Preparation Apparatus prior to heat treatment may be held at 2-27 C for up to 24 hours prior to testing.
4.    Remove the tip cap from the Sample Preparation Apparatus and squeeze five to ten drops of sample into a clean Alethia Heat Treatment Tube.
5.
Repeat Sample Preparation Steps for all samples to be processed. 
6. 
Heat the Alethia Heat Treatment Tubes containing mixed sample in a dry-bath/heat block at 95 C +/- 5 C for 10 +/- 2 minutes.  Use the digital thermometer and the interval timer to monitor temperature and time. 
7.  Remove the tubes from the dry-bath/heat block and vortex for approximately 10 seconds. Heat treated Samples may be held at 2-8 C for 4 hours OR may be frozen and stored a         ≤-20 C for 1 day prior to addition to Reaction Buffer.  Heat treated samples may be frozen and thawed once.
TEST PROCEDURE:
NOTE: A maximum of 10 samples can be processed in a single Alethia reader run. 
1. Transfer 50 µL of heat-treated sample to an appropriately labeled Alethia Reaction Buffer tube.

2. Vortex the Reaction Buffer Tube containing heat-treated sample for approximately 10 seconds.

3. Repeat steps 1 and 2 for all the samples to be analyzed before proceeding.

4. Remove 1 Alethia C. difficile Test Device from its protective pouch per sample.  Carefully open the device, holding the tubes such that the lyophilized reagent will not fall out upon opening.  Place device on a flat surface or in a rack that can accommodate the device. 

5. Using a micropipette, first transfer 50 µL from the Reaction Buffer tube containing the heat treated sample to the TEST chamber (left/White Bead), and then transfer 50 µL from the Reaction Buffer tube containing the heat treated sample to the CONTROL chamber (right/Yellow Bead) of the Alethia Test Device.  Take care not to introduce air to the reaction mixture.  Do not mix reactions with pipette.  Close the Alethia Test Device and fasten the latch securely. 

6. Gently mix the device by tapping or flicking, being careful to not introduce air bubbles or get the reaction mixture into the device lid.  Carefully examine the reaction tubes for air bubbles or liquid in the lid.  If air bubbles or liquid is noted, tap the device on the bench top and repeat the visual inspection.

7. Insert the Alethia Test Device into the Alethia reader and initiate amplification reaction and detection.  Results will be displayed at the conclusion of the run.
CALCULATIONS:

There are no calculations associated with this procedure.
INTERPRETATION OF RESULTS:
	Sample ID 
	Reported Result
	Interpretation

	Patient Specimen
	POSITIVE
	Sample contains toxigenic C. difficile strain with the pathogenicity locus (PaLoc).

	
	NEGATIVE
	No toxigenic C. difficile detected.

	
	INVALID
	No reportable result.  Repeat the test using the original stool sample.
Inhibitory patient specimen, improper sample preparation, reagent failure, instrument failure or internal control failure. 

	Positive Control
	POSITIVE
	Valid positive control result.  Reagents active at time of use, Alethia reader performing correctly.

	
	NEGATIVE
	Incorrect control result.  Repeat control testing as the first step in determining the root cause of the failure.  If control failures are repeated please contact Meridian’s Technical Services at 1-800-343-3858 (US) or your local distributor.

	
	INVALID
	No reportable result.  Repeat entire assay run using original stool samples. 
Improper sample preparation, reagent failure, instrument failure or internal control failure.

	Negative Control
	POSITIVE
	Incorrect control result.  Repeat control testing as the first step in determining the root cause of the failure.  If control failures are repeated please contact Meridian’s Technical Services at 1-800-343-3858 (US) or your local distributor.

	
	NEGATIVE
	Valid negative control result.  Reagents active at time of use, Alethia reader performing correctly.

	
	INVALID
	No reportable result.  Repeat entire assay run using original stool samples. 
Improper sample preparation, reagent failure, instrument failure or internal control failure.

	EMPTY WELL
	NONE
	No Alethia Test Device in the Alethia reader Well.
OR
The Alethia Test Device present is compromised due to sample preparation failure, dirty device or improperly seated device.  Repeat the test using original sample.


REPORTING OF RESULTS: 
Positive Test: Sample contains toxigenic C. difficile strain with the pathogenicity locus (PaLoc)
Negative Test: No toxigenic C. difficile detected or is below the limit of detection.
EXPECTED VALUES:
The frequency of antibiotic-associated diarrhea caused by C. difficile is dependent on several factors, including patient population, type of institution, and epidemiology.  The incidence of CDAD in patients suspected of having antibiotic-associated disease is 15-20%, although different facilities may find positive rates above or below this range.  The incidence of toxigenic C. difficile during the 2010 period of this study was 15.5% and 16.1%. 
ASSAY REACTIVITY:

The following C. difficile stock cultures from different sources were tested and produced positive reactions at 64 CFU/test with Alethia C. difficile: 

Type 0:  Strains 10463, 2004111, 2004205, 2005070, 2005257, 2008029, 2008162, 2008341, 2008351, 2009066, 2009099, B1, G1, J7, K12, Y1; Type III: 2004052, 2004118, 2007431, BI17, BI8; Type V: 2005325, 2006240, 2008188, 2009018, 2009065, BK6; Type VIII: 43598, 2008016, CF1; Type X 8864; Type XII 2007435; Type IX/XXIII 2007858; Unknown Type 2009132, 2009155, 2009277.

CROSSREACTIVITY:

Crossreactivity studies were performed with positive and negative stool specimens inoculated with bacterial or fungal organisms to a final concentration of 1.2 x 108 or virus at a minimum of 1 x 105.06 TCID50/mL.  None of the following organisms in stool reacted with Alethia C. difficile:

Aeromonas hydrophila, Bacteroides fragilis, Campylobacter coli, Campylobacter fetus, Campylobacter jejuni, Candida albicans, Citrobacter freundii, Clostridium sordellii, Clostridium perfringens, Enterobacter cloacae, Enterococcus faecalis, Escherichia coli, Escherichia coli O157:H7, Escherichia fergusonii, Escherichia hermannii, Helicobacter pylori, Klebsiella pneumoniae, Lactococcus lactis, Listeria monocytogenes, Peptostreptococcus anaerobius, Plesiomonas shigelloides, Proteus vulgaris, Pseudomonas aeruginosa, Pseudomonas fluorescens, Salmonella Groups B-E, Serratia liquefaciens, Serratia marcescens, Shigella boydii, Shigella flexneri, Shigella sonnei, Staphylococcus aureus, Staphylococcus epidermidis, Vibrio parahaemolyticus, Yersinia enterocolitica, Adenovirus Types 40 and 41, Coxsackievirus, Echovirus, Rotavirus.

TEST FOR INTERFERING SUBSTANCES:

The following substances, at the specified saturated solvent/diluents concentrations, do not interfere with test results in the final concentrations listed:  Barium sulfate (5 mg/mL), fecal fat (equivalent to 2.65 mg stearic plus 1.3 mg palmitic acids per mL), hemoglobin (as methemoglobin) (3.2 mg/mL), IgA (5 mg/mL), Imodium AD® (0.00667 mg/mL), Kaopectate® (0.87 mg/mL), Metronidazole (12.5 mg/mL), mucin (3.33 mg/mL) Mylanta® (4.2 mg/mL), Pepto-Bismol® (0.87 mg/mL), Prilosec® (0.5 mg/mL), Tagamet® (0.5 mg/mL), TUMS® (0.5 mg/mL), Vancomycin (12.5 mg/mL), white blood cells (5% V/V), whole blood (5% V/V).

LIMITATIONS OF THE PROCEDURE:
1. Assay does not distinguish between viable and non-viable organisms.
2. Test results are to be used in conjunction with information available from the patient clinical evaluation and other diagnostic procedures. 
3. This test is for use with unformed stools preserved or unpreserved stools only.  Performance characteristics of other clinical specimen types have not been established.
4. Two distinct groups have been identified that can harbor C. difficile asymptomatically at very high rates.  Colonization at rates up to 50% and higher have been reported in infants and rates up to 32% in cystic fibrosis patients.
5. This assay does not identify antimicrobial susceptibility.
6. This test detects but does not differentiate the NAP1 (Ribotype 027) strain from other toxigenic strains of C. difficile.
7. The detection of bacterial nucleic acid is dependent on proper specimen collection, handling (including transportation and storage) and preparation (dilution and heat treatment).  Failure to follow instructions for collection, handling, and preparation may cause incorrect results.
PERFORMANCE CHARECTERISTICS:

Refer to Directional Insert- Meridian Bioscience Alethia C. difficile
REFERENCES:
Refer to Directional Insert- Meridian Bioscience Alethia C. difficile
Meridian Bioscience

3471 River Hills Drive

Cincinnati, OH 45244 USA



Ph:  (800) 343-3858, (513) 271-3700

SN11016 REV 01/20

