[image: image1.png]meridian BIOSCIENCE
S





Page 7 of 7

Procedure and revision number: SN11240   CLSI 12/16
Procedure:  ImmunoCard STAT Rotavirus

Meridian Catalog 750030
	Institution: 

	Address:

	Department:


	Prepared By:
	Date Adopted
	Supersede procedure #:

	
	
	

	
	
	

	
	
	

	
	
	

	
	
	


	Distributed To:
	# of Copies
	Distributed To:
	# of Copies

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	

	
	
	
	


NOTE: This procedure is provided to Meridian’s customers to assist with the development of laboratory procedures.  This document was derived from, and was current with, the instructions for use (IFU) that accompanies the product at the time it was created. The user is instructed to consult the IFU packaged with the product to ensure currency of the procedure prior to adapting the document to routine laboratory use and periodically thereafter to ensure future IFU modifications, which might effect this procedure, are identified.   Any modifications to this document are the sole responsibility of the person making the modifications.

PRINCIPLE:

The ImmunoCard STAT! Rotavirus Immunoassay is a rapid in vitro qualitative procedure for the detection of rotavirus antigen in human stool.  The test can be used to aid in the diagnosis of rotavirus associated gastroenteritis.

The ImmunoCard STAT! Rotavirus assay detects the presence of rotavirus antigen in stool.  Patient specimen is diluted 1:15 in Sample Diluent.  The suspension is mixed and 150 µL is added to the sample port of the device.  The sample mobilizes gold particles coated with monoclonal antibody to rotavirus and migrates along the membrane through the Test (polyclonal anti-rotavirus antibody) and Control zones.  After ten minutes, the Test and Control zones are observed for the presence of red/purple lines across the membrane surface.  If rotavirus is present in the sample, a complex is formed between the capture antibody and the monoclonal antibody-gold conjugate which can be seen visually as a red/purple line in the Test zone.  No red/purple line in the Test zone indicates a negative result. The Control line serves as a procedural control to assure that the sample has migrated the appropriate distance along the membrane.

SPECIMEN:

Preferred sample types:

Solid, semi-solid or liquid samples are acceptable. Samples obtained from a diaper are acceptable. 
Undesirable specimens:

1. Samples in transport media.
2. Samples with detergent residue.
3. Samples that have dried out.
Collection:
Stool specimens may be collected by the method routinely utilized by the laboratory, providing no dilution of the specimen occurs.  A sample may be collected into a clean, dry container free of detergent residue, or obtained from a diaper.
The specimen should be tested as soon as possible, but may be stored up to 72 hours at 
2-8 C prior to testing.  If testing cannot be performed within this time frame, specimen should be frozen in a non-defrosting freezer (-20 C or lower).
For the best results, specimens should be collected after onset of symptoms.  Several authors have reported declining numbers of rotavirus particles after day eight or nine, with peak counts occurring on days three through five.  Samples collected after day eight or nine may be less reactive than those collected earlier in the course of the disease.  Specimens containing high levels of blood may fail to flow in the ImmunoCard STAT! Rotavirus device, resulting in an invalid test result.  Testing of an additional specimen is recommended under such circumstances.
This facility’s procedure for specimen collection is: _________________________________ __________________________________________________________________________
This facility’s procedure for transporting specimens is: ______________________________

 __________________________________________________________________________
This facility’s procedure for rejected specimens is: _________________________________ __________________________________________________________________________
MATERIALS AND EQUIPMENT:

Materials:

1. ImmunoCard STAT! Rotavirus Devices - Individually foil pouched devices containing immobilized rabbit anti-rotavirus antibody (Test zone), goat anti-mouse IgG antibody (Control zone) and embedded monoclonal antibody to rotavirus VP-6 conjugated to gold particles.

2. Positive Control -Inactivated rotavirus (SA-11) in a buffer containing 0.02% thimerosal as a preservative.

3. Sample Diluent - Buffer containing 0.1% sodium azide as a preservative.

4. Transfer Pipettes 

Materials Not Provided:

1. 12 X 75 mm Test tubes
2. Applicator Sticks
3. Interval timer
Preparation:

Bring specimens and reagents to room temperature (21 - 25 C) before testing. 

PERFORMANCE CONSIDERATIONS:

1. All reagents are for in vitro diagnostic use only.

2. Reagent concentration, incubation times and temperatures (21-25 C) have been optimized for sensitivity and specificity.  Best results are obtained by adhering to these specifications.  Once the assay has been started, complete all subsequent steps without interruption.

3. Patient specimens and used ImmunoCard STAT! Rotavirus devices may contain infectious agents and should be handled at Biosafety Level 2 as recommended in the CDC/NIH manual “Biosafety in Microbiology and Biomedical Laboratories”.

4. Positive Control reagent contains inactivated rotavirus.  However, it should be handled as a potential biohazard.

5. All reagents should be gently mixed and at 21-25 C before use.

6. Do not interchange reagents from different kit lot numbers, or use expired reagents.

7. Hold Positive Control vial vertically to insure proper drop size and delivery.  Do not allow the tips of the vial or pipette to touch the sample port.

8. Use only one transfer pipette per control or specimen.  Discard after use.  Do not attempt to reuse.

9. Stool must be mixed thoroughly (regardless of consistency) to insure a representative sample prior to pipetting.  Do not use stools that have dried out.

10. Dilution of stool as described under Sample Dilution is important.  Over-inoculation of stool into the Sample Diluent may restrict movement within the ImmunoCard STAT! Rotavirus device so as to produce an invalid result.

STORAGE REQUIREMENTS:
Store kit refrigerated 2 - 8 C and return kit to the refrigerator promptly after each use.  DO NOT FREEZE.
At this facility, kits are stored:_______________________________________________ ​​​​​                                                                                                           
CALIBRATION:

Not applicable to this assay.
QUALITY CONTROL:
The Positive and Negative Controls should be assayed once upon receipt of the kit.  Thereafter, Meridian recommends that the procedural control, performed with each test, is sufficient to verify reagent integrity and assay performance.  The negative and positive controls are intended to monitor for reagent failure, but will not ensure precision at the analytical assay cut-off.  Additional controls may be tested in accordance with guidelines or requirements of local, state, and /or federal regulations and accrediting organizations.

Add three drops of Positive Control or, using a transfer pipette, add 150 µL Sample Diluent (Negative Control) directly to Sample port of appropriate device (do not dilute Positive Control).

1. The Positive Control should yield visually detectable red/purple Test and Control lines.

2. The Negative Control should yield a visually detectable red/purple Control line.  No Test line should be present.

QC Testing Frequency and Documentation:

For this facility, External QC is run:______________________________________________                                                                                                   

Results of External QC and action(s) taken when control results are unacceptable are documented:________________________________________________________________

PROCEDURE:
1. Sample Dilution

a. Add 350 µL Sample Diluent to one 12 x 75 mm test tube for each specimen to be tested.

b. Mix stool thoroughly, regardless of consistency.

c. Liquid or semi-solid stool: Using a transfer pipette, draw stool to the 25 µL calibration point (first mark from tip of the pipette).  Dispense the stool into the Sample Diluent in appropriate 12 x 75 mm tube.  Using the same pipette, gently withdraw and expel the stool suspension several times, then vortex ten seconds.  Leave transfer pipette in tube for further use.  Note: Do not pipette more than 25 µL of stool.  Over-inoculation with stool may produce invalid results.  Proceed to Step 2 within 30 minutes.

d. Solid stool (nonpipetable stool from a diaper): Using a wooden applicator stick, transfer a 2 mm diameter portion of stool into the Sample Diluent in the appropriate 12 x 75 mm tube.  Emulsify the stool thoroughly using the applicator stick, then vortex ten seconds.  Place transfer pipette in the tube.  Proceed to Step 2 within 30 minutes.

2. Remove appropriate number of ImmunoCard STAT! Rotavirus devices from their pouches.  Label appropriately.  Use one device per control or sample.

3. Vortex each diluted specimen for ten seconds.  Using the original specimen transfer pipette, draw diluted sample to the 150 µL calibration point (second mark from tip of the pipette) and add to Sample port.

4. Incubate ten minutes at 21-25 C.  NOTE: During the ten minute incubation, diluted specimen must move past the Control zone.

5. Visually read Control and Test zones for the presence or absence of a red/purple line at the end of the incubation period.

This facility’s procedure for patient preparation is:___________________________  __________________________________________________________________

This facility’s procedure for sample labeling is:_______________________________ ____________________________________________________________________ 

CALCULATIONS:

There are no calculations associated with this procedure.

INTERPRETATION OF RESULTS:

Positive Test Result: Visually detectable red/purple Test and Control lines.  A positive result indicates the presence of rotavirus antigen.
Negative Test Results: Visually detectable red/purple Control line.  No red/purple Test line present.  A negative result indicates that rotavirus antigen is absent or below the level of detection.
Invalid Test Result: No visually detectable red/purple Control line, with or without a visually detectable red/purple Test line.
In the event this test becomes inoperable, this facility’s course of action for patient samples is: ______________________________________________________________

_______________________________________________________________________
REPORTING OF RESULTS: 
Negative for Rotavirus: Report test result as “Rotavirus antigen is absent or below the level of detection.”  

Positive for Rotavirus: Report test result as “Rotavirus antigen detected.”
This facility’s procedure for patient result reporting is:__________________________________ ____________________________________________________________________________________________________________________________________________________________                                                                         
EXPECTED VALUES:

The ImmunoCard STAT! Rotavirus test detects the presence of rotavirus antigen in stool.  Expected values for a given population should be determined for each laboratory.  The rate of positivity may vary depending on patient age, geographic location, season, method of specimen collection, handling and transportation, and a general health environment of the patient population under study.
It has been reported that in neonates, when rotavirus was present, the disease was mild or totally asymptomatic.  However, during cooler months, rotavirus may account for approximately 50% or more of the gastroenteritis found in hospitalized children.  In adults, the incidence of serious gastroenteritis caused by the virus is relatively low and when infected, adults tend to be asymptomatic.  Studies from nursing homes and hospital geriatric wards show that this population is at an increased risk and susceptible to rotavirus associated disease.
LIMITATIONS OF THE PROCEDURE:
1. The ImmunoCard STAT! Rotavirus test does not define the presence of rotavirus associated gastroenteritis, but only demonstrates the presence of the antigen in stool.  As with all in vitro diagnostic procedures, test results should be interpreted by a physician in conjunction with other clinical information.

2. Limit of detection in stool specimens is 1.8 - 3.7 x 106 rotavirus particles (SA-11) per test volume.

3. The use of meconium stools in this assay is not recommended as their performance characteristics have not been evaluated.

4. A positive result does not preclude the presence of other infective organisms.

ASSAY SPECIFICITY:
The ImmunoCard STAT! Rotavirus test was evaluated for specificity using the reference bacterial and viral strains listed below.  A stool positive for rotavirus and a stool negative for rotavirus were spiked with bacteria (≥ 6 X 108 cfu / mL stool) or viruses (≥ 1 X 102.7 TCID50/mL).

Microorganism or Virus (# strains tested)

	Aeromonas hydrophilia (1)
	Escherichia coli (O157:H7) (1)
	Staphylococcus aureus (1)

	Bacillus cereus (1)
	Escherichia coli 
(non-O157:H7, SLT+) (1)
	Staphylococcus aureus 
(Cowan 1) (1)

	Bacillus subtilis (1)
	Helicobacter pylori (1)
	Streptococcus faecalis (1)

	Bacteroides fragilis (1)
	Klebsiella pneumoniae (1)
	Vibrio cholerae (1)

	Campylobacter coli (1)
	Proteus mirabilis (1)
	Vibrio parahaemolyticus (1)

	Campylobacter fetus (1)
	Pseudomonas aeruginosa (1)
	Adenovirus, types 2, 40, 41

	Campylobacter jejuni (1)
	Salmonella (1)
	Coronavirus

	Citrobacter freundii (1)
	Serratia liquefaciens (1)
	Coxsackie types A9, B1, B6

	Clostridium difficile (1)
	Shigella dysenteriae (1)
	Echovirus, types 22, 32

	Clostridium sordellii (1)
	Shigella flexneri (1)
	Enterovirus, type 69

	Enterobacter aerogenes (1)
	Shigella sonnei (1)
	Poliovirus, type 1

	Escherichia coli (1)
	
	


All organisms were found to be negative when spiked into the negative stool.  In addition, they did not interfere with either the positive specimen or the procedural control.

Studies with different serotypes of group A rotavirus (the most common form) showed that the ImmunoCard STAT! ® Rotavirus test detects serotypes 1, 2, 3, and 4.  Reactivity with serotypes 8 and 9 has not been evaluated.

PERFORMANCE CHARACTERISTICS:
Refer to Directional Insert – Meridian Bioscience ImmunoCard STAT! Rotavirus
REFERENCES:
Refer to Directional Insert – Meridian Bioscience ImmunoCard STAT! Rotavirus
Meridian Bioscience

3471 River Hills Drive

Cincinnati, OH 45244 USA



Ph:  (800) 343-3858, (513) 271-3700

SN11240  CLSI 12/16

